## Slide 1
Developing 2D and 3D Micropolar FEM Models for Porous GBR Meshes in Dentistry Applications
22-24 October 2024, Florence, Italy
AbdolMajid Rezaei
Prof. N. Fantuzzi
Dr. Razie Izadi

## Slide 2
Introduction
Guided Bone Regeneration (GBR) Mesh
Using a mechanical barrier, such as a membrane, to separate and protect the area of bone loss from the surrounding tissue.
1 Exposing the Bone Defect: A small incision to access the area of defect.
2 Bone Grafting: a cement that may contain bone, antimicrobial additives and stimulants that promotes new bone growth is placed under the membrane.
3 Preparing GBR Mesh: cut and formed by the surgeon
4 Placing and fixing the GBR Mesh
Bone Defect
1
2
3
4
GBR Mesh
https://www.indigoortho.co.uk/gbr-systems
Porosity
Stiffness
New bone can grow in the designated area.
Enough stiffness to create and maintain a suitable space for the intended bone regeneration.
Enough porosity to facilitate the diffusion of fluids, oxygen, nutrients, and bioactive substances for cell growth.

## Slide 3
Introduction
Equivalent Homogenised Models for Porous Structures
When the geometry includes pores, there is an internal structure.
By using equivalent model and defining effective properties instead of the real microstructure, we can decrease computational expense of modelling and discretisation.
Non-classical continuum theories can consider the internal structure without direct modelling of details.
In the present work, non-classical micropolar theory as the homogenised equivalent model is implemented for the porous structures in 2D and 3D.
  Internal Structure
  Equivalent Continuum Model
?
  Porous Geometry

Non-classical Theories
Micropolar Continua
Micromorphic Continua

Second-Gradient Continua
…

## Slide 4
Introduction
Micropolar Theory
 | Micro-rotation
 | Nonsymmetric Strain
 | Curvature
 | Nonsymmetric Stress
 | Couple-stress
 | Body Force
 | Body Couple
 | Micro-rotation
 | Nonsymmetric Strain
 | Curvature
 | Nonsymmetric Stress
 | Couple-stress
 | Body Force
 | Body Couple
New strain measure as Curvature.
New stress measure as Couple-stress.
Degrees of Freedom: U,
Kinematic Relations:
Balance Equations:
Constitutive Equations: (Linear Elastic)

## Slide 5
Methodology
2D Tetragonal Micropolar Model for Porous Plate
A
D
A
  5 Independent Material Parameters
General linear elastic constitutive equations in plane
Ortho-tetragonal Material Symmetry
Considering the linear elastic micropolar in 2D, for the geometries we study here, 5 independent material parameters are required to define the equivalent model.

## Slide 6
Methodology
Homogenization of 2D Model
The heterogeneous structure, is called Micro Model and is described by Cauchy continuum.
Homogenised model is called Macro Model and is described by constitutive parameters of a micropolar continuum.
Kinematic Map: Links the two levels of description
Express the microscopic displacement field within the RVE as a function of the macroscopic strain measures at the material point on the macro-level.
Micro Model
Cauchy
u,v
Macro Model
Micropolar
Equivalent
Model
U,V,Φ

## Slide 7
Methodology
Homogenization of 2D Model
Once the kinematic map is determined, we proceed to discover the micropolar material parameters.
We use the finite element method (FEM) to compute the response of the porous structure under different loadings.
  Extract the total elastic strain energy stored in the RVE from COMSOL.
  Equating to the energy of an equivalent micropolar continuum from analytical expression.
For each scenario, the micropolar material parameters are determined in such a way that the homogenized model retains the same amount of strain energy when exposed to the same loading conditions.

## Slide 8
Methodology
Finding Effective Micropolar Material Parameters for 2D Model
Micropolar Model
Pore
Cauchy Continua
Bending Test
Biaxial Test
Uniaxial Test
Shear (SYM) Test
Rotation Test
Numerical FEM Simulation

## Slide 9
Methodology
Extending the Homogenization Procedure for FG Porous Structures
Equivalent homogenized model for Functionally Graded (FG) porous structure is derived by considering the homogenization procedure developed for homogenous porosities.

## Slide 10
Methodology
Extending the Homogenization Procedure for FG Porous Structures
First, a parametric study is conducted to find the equivalent parameters of uniform porous plates with various pore sizes.
In this parametric study, the pore density (the number of pores per unit length) is kept constant.
...
A1111, A1122
A1212,A1221
D11
A1111, A1122
A1212,A1221
D11
A1111, A1122
A1212,A1221
D11
...

## Slide 11
Methodology
Extending the Homogenization Procedure for FG Porous Structures
Pore sizes are changed to find the required equivalent parameters for each section of the FG porous structure.
...
A1111, A1122
A1212,A1221
D11
Chart: Pore Size (𝒍 𝒑)
| 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 6.9999999999999896E-2 | 0.08 | 0.09
A1111 | 4.0765 | 3.8615 | 3.5418 | 3.1581 | 2.7451 | 2.327 | 1.9165 | 1.5165 | 1.1224
A1111, A1122
A1212,A1221
D11
A1111, A1122
A1212,A1221
D11
...

## Slide 12
Methodology
Extending the Homogenization Procedure for FG Porous Structures
These data are then utilised to find curve-fitted functions for each material parameter with respect to the pore size.
...
A1111, A1122
A1212,A1221
D11
Chart: Pore Size (𝒍 𝒑)
| 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 6.9999999999999896E-2 | 0.08 | 0.09
A1111 | 4.0765 | 3.8615 | 3.5418 | 3.1581 | 2.7451 | 2.327 | 1.9165 | 1.5165 | 1.1224
Chart: Pore Size (𝒍𝒑)
| 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 6.9999999999999896E-2 | 0.08 | 0.09
A1111 | 4.0765 | 3.8615 | 3.5418 | 3.1581 | 2.7451 | 2.327 | 1.9165 | 1.5165 | 1.1224
…
A1111, A1122
A1212,A1221
D11
A1111, A1122
A1212,A1221
D11
...

## Slide 13
Methodology
Homogenization of 3D Model
A size-effect is predicted by micropolar theory for the torsion of circular cylinders.
It is possible to determine the micropolar parameters by measuring the torsional rigidity vs size.
Specimens at different length scales, while keeping unit cell size constant.
d: Diameter
J: Rigidity
Cauchy Detailed Model
Micropolar
Equivalent Model

## Slide 14
Methodology
Homogenization of 3D Model
Bio-inspired 3D microstructures such as gyroids and other Triply Periodic Minimal Surfaces (TPMS) can be employed for designing GBR bone scaffolds.
Studying the impact of changes in microstructural architecture which reflects in the equivalent micropolar parameters.
The unit cell design, porous structure and CAD model are developed using nTop software.
Final 3D meshed part is imported in COMSOL for further FEM analyses.
CAD Model
FEM Tests

## Slide 15
Multiplying balance equations by the test functions, integrating over computational domain D
Based on the product rule of derivatives:
  Strong Form
Balance Equations
  Weak Form
  Test Function for
  Test Function for U
Methodology
FEM Micropolar Model in COMSOL Using PDE Weak Form
And by using the divergence theorem:

## Slide 16
Methodology
FEM Micropolar Model in COMSOL Using PDE Weak Form
By using partial differential equation (PDE) modelling in COMSOL instead of traditional FE modelling, no user subroutine is required.
Various complex geometries, B.C., and loadings can be applied in a user-friendly graphical interface.
Visualisation of the results is convenient.
-(s11*test(ux)+s12*test(uy)+s21*test(vx)+s22*test(vy) +s23*test(vz)+s31*test(wx)+s32*test(wy)+s33*test(wz)) + p1*test(u)+p2*test(v)+p3*test(w)
-(mu11*test(phi1x)+mu12*test(phi1y)+mu13*test(phi1z) +mu21*test(phi2x)+mu22*test(phi2y)+mu23*test(phi2z) +mu31*test(phi3x)+mu32*test(phi3y) + mu33*test(phi3z)) -((s12-s21)*test(phi3)+(s31-s13)*test(phi2) +(s23-s32)*test(phi1))+q1*test(phi1)+q2*test(phi2) +q3*test(phi3)

## Slide 17
Methodology
FEM Micropolar Model in COMSOL / Benchmarks for 2D Model
Patch Tests
Test 1 (Cauchy)
Test 2 (Micropolar, without Curvature)
Test 3 (Micropolar + Curvature)
E. Providas, M.A. Kattis, 2002
DOI:
10.1016/S0045-7949(02)00262-6
 | @ | Benchmark | COMSOL
u1 | P | 0.1950 | 0.19500
u2 | P | 0.2100 | 0.21000
phi | P | 0.0004 | 0.00040
S11 | Node 2 | 4.0000 | 4.00000
S21 | Node 2 | 1.4667 | 1.46670
Mu31 | Node 2 | 0.0400 | 0.04000
 | L / a | Element | Benchmark | COMSOL
1 | 3.0 | 4Node | 2.49640 | 2.50110
2 | 3.0 | 9Node | 2.42046 | 2.42128
3 | 10.0 | 4Node | 2.28474 | 2.29980
4 | 10.0 | 9Node | 2.22525 | 2.26834
M. Tuna et al., 2021
DOI:
10.1007/978-3-030-63050-8_11
Stress Concentration Factor
Micropolar plate with a circular hole under uniform tension.
Stress concentration factor (SCF) is compared with those of Reference.

## Slide 18
Methodology
FEM Micropolar Model in COMSOL / Benchmarks for 3D Model
Force Patch Test
Cantilever beam subject to pure tension
Seeking the state of constant stress.
For an arbitrary number of finite elements.
S. Grbčić et al., 2018
DOI:10.1016/j.compstruc.2018.04.005
Displacement Patch Tests
A generalization of 2D patch tests.
Discretization with 7 arbitrarily distorted hexahedral finite elements.
 | Benchmark | COMSOL
u1() | 0.600 | 0.60000
u2() | 0.700 | 0.70000
u3() | 0.700 | 0.70000
phi1 () | 2.100 | 2.10000
S11 | 5.000 | 5.00000
Mu11 | 0.020 | 0.02000
Mu12 | -0.0400 | -0.04000
 | Benchmark | COMSOL
 | 0.600 | 0.60000
 | 0.700 | 0.70000
 | 0.700 | 0.70000
 | 2.100 | 2.10000
S11 | 5.000 | 5.00000
Mu11 | 0.020 | 0.02000
Mu12 | -0.0400 | -0.04000

## Slide 19
Results
Parametric Study for Various Pore Patterns Using COMSOL
  Pore Type & Pattern:
  Rectangular pores
  Honeycomb pattern
  Triangular pores, Honeycomb pattern
  ...
  Pore Size:
  Different length scale
  Pore Density:
  Different No. of pores per length

## Slide 20
Results
Parametric Study for Various Pore Patterns Using COMSOL / Parametrisation
Titanium Alloy Ti6Al4V | Poly- Lactic Acid (PLA)
Young’s Modulus, GPa | 
114 | 1.03
Poisson Ratio | 
0.30 | 0.32
Density, kg/m3 | 
4500 | 1250
    Mechanical Characterization of 3DPrinted Individualized Ti-Mesh for Alveolar Bone Defects L. Bai et al.
    Physical and mechanical properties of PLA S. Farah et al.

## Slide 21
Results
Parametric Study for Various Pore Patterns/Materials Using COMSOL
Ti6Al4V
PLA

## Slide 22
Discussion
Using the Equivalent Mechanical Parameters for Designing GBR Mesh
Optimum Design for Porous GBR Meshes:
Nearest mechanical properties to the bone.
GBR mesh is in contact with the cortical (compact) bone.
Try to make the material parameters of the GBR mesh consistent with its adjacent bone.
Experimental estimate of the micropolar parameters of compact bone.
Finding a configuration with the material parameters close to those reported for compact bone in the literature.
    Experimental study of micropolar and couple stress elasticity in compact bone in bending J. Yang & R. Lakes
 | Value | Unit
A1111 | 12.00 ~ 43.43 | GPa
A1122 | 4.00 | GPa
A1212 | 21.10 ~ 36.77 | GPa
A1221 | -13.05 ~ 2.67 | GPa
D11 | 3.24 | kN

## Slide 23
Discussion
Using the Equivalent Mechanical Parameters for Designing GBR Mesh
Circular pores, Titanium Alloy
 Parameters are beyond the required values.
Rectangular pores, Titanium Alloy
 Pore sizes of 0.13 L - 0.15 L and porosity about 0.7, a good agreement for A1122, A1212, A1221 can be achieved.
 | Value | Unit
A1111 | 12.00 ~ 43.43 | GPa
A1122 | 4.00 | GPa
A1212 | 21.10 ~ 36.77 | GPa
A1221 | -13.05 ~ 2.67 | GPa
D11 | 3.24 | kN

## Slide 24
Discussion
Functionally Graded (FG) Porous Design for GBR Mesh
GBR meshes are fixed to the mandible bone using biocompatible screws
At these fixing locations higher stiffness is required. Therefore, smaller pore sizes are more desirable there.
Mimicking natural FG structure of the bone
FG structure of type O is suggested
Central part  mechanical properties close to cancellous (trabecular) bone while providing a proper diffusion properties.
Part near fixing areas  as near as possible to cortical (compact) bone to provide required load-bearing capacities.
Conventional Design (uniform porosity)
FG Porous Design

## Slide 25
Currently, metallic GBR sheets are used as a mechanical barrier and bone grafts are inserted into the space created by the barrier.
After the healing period, it is necessary another surgery for removing the GBR mesh.
To avoid a second surgery, a recent solution is to use biodegradable materials such as polylactic acid (PLA) instead of metallic meshes.
However, suboptimal mechanical properties of biodegradable materials such as low stiffness and strength have limited their application.
Discussion
3D FG Design for Porous GBR Dental Scaffold
GBR Mesh
(Space Maintaining)
Bone Grafts
(Stimulating Bone Regeneration)
3D GBR Dental Scaffold

## Slide 26
With the developed framework, an innovative 3D design for GBR meshes is suggested.
The microstructure evolves from the porous structure at the lower level to the compact structure at the top surface.
The porous structure can host bio-active agents for stimulating bone regeneration
While the upper compact surface should provide the required stiffness.
By implementing our model, the distribution in both the internal structure and the material properties (for instance using bimodal material such as PLA+HA (Hydroxy appetite) or adding nano-reinforcements) can be customized.
Discussion
3D FG Design for Porous GBR Dental Scaffold
3D Design
for
GBR Dental Scaffold
Porous
Dense

## Slide 27
  For Your Attention and Time…

## Slide 28
Questions

## Slide 29
Results
Effectiveness of Homogenized Micropolar Model
Micro Model
Macro Model
Comparison of displacement magnitude contours for detailed and homogenized micropolar model.
Indentation Test.
Predicted values are close to those of the detailed model